## Slide 1
INTERNATIONAL DEPLOYMENT OF ACCELEROMETERS (IDA)
Carl Ebeling, IGPP, SIO, University of California, San Diego
Observatorio Vulcanológico y Sismológico de Costa Rica
(OVSICORI), Heredia, Costa Rica, 30 January 2019
Planned Work at
Global Seismograph Network (GSN)
Seismic Station JTS
(Las Juntas de Abangares)

## Slide 2
(41)
(~100)
(12)

## Slide 3
GSN requirement: Record on scale all seismic waves from Mw=9 earthquake at 3000 km

## Slide 4
Primary and Secondary Sensors at GSN Vault Stations
UOSS (United Arab Emirates)
STS-1s and STS-2
JTS: STS-1s and STS-2 (not shown)

## Slide 5
Aging STS-1 Seismometers
JTS: STS-1s and STS-2 (not shown)
Streckeisen STS-1Z/VBB and STS-1H/VBB seismometers are experiencing increasing problems everywhere due to age.
At prime, exceptional performance. But now:
No longer manufactured
Spare parts not available
Replacements only via cannabalism
Upkeep and maintenance in turn more challenging

## Slide 6
STS-1 Seismometers: Aging at JTS, too
2018-026 12:32 M5.6 13 km WNW of Santa Maria, Colombia
(delta = 11.7 deg.)
30-minute-long record
4-hour-long record
STS-1 V
STS-1 H1
STS-1 H2
STS-2.5 Z
STS-2.5 H1
STS-2.5 H2

## Slide 7
LP (360 s) Replacements for STS-1 Seismometers
Few low-noise very long-period (360 s)seismometers:
Heavy investment in R & D
Testing difficult
Significant engineering and manufacturing challenges
High cost but low demand
Solution:
Install borehole-package STS-6A in cored hole in vault
Sensitivity: 2 x 600 V/m/s
Response: Flat to ground velocity, 2.78 mHz to 50 Hz
Package: 153 mm diameter x 610 mm long (6 in dia x 24 in)
Weight: 30 kg
Power draw: 0.9 W, max < 2 W
STS-6A

## Slide 8
Core Drilling at JTS
Coring rig
Core extraction tool
Goal depth ~ 1m
~5-6 cm/h

## Slide 9
Determination of Seismometer Orientation
Elements of independent calibration/recording system are used to determine orientation of seismometer
Nanometrics Compact Trillium 120 and Taurus digitizer

## Slide 10
Determination of Seismometer Orientation
APS used to determine orientation of Compact Trillium

## Slide 11
Determination of Seismometer Orientation
Determination of relative ampitude accurate to < 1 %
Determination of relative orientation accurate to < 1 deg
2+ hours of microseismic background noise usually sufficient for robust analysis

## Slide 12
Thank you!
Pier at the Scripps Institution of Oceanography, UCSD

## Slide 13
Pier at the Scripps Institution of Oceanography, UCSD
Thank you!